GER Wave-Mixed Bioreactor

GER series wave bioreactor is an efficient cell culture system,with optional sensor modules and the automated control software system, and paired with the GER single-use wave-mixed bioreactor bags. The working volume can reach 25L, providing solutions for general life science research in the biopharmaceutical field, seed tank cultivation, and cGMP production.

Excellent Cultivation Performance

• Rapid mixing and improvement of gas mass transfer;

• No need of bubble ventilation during oxygen mass transfer to avoid cell damage and the use of defoaming agent;

• Improving cell status and enhancing density and yield;

Flexible Configuration and Convenient Operation

• Interchangeable tray design, compatible with 10L, 20L, and 50L single-use bags.

• Reduction of vessel contents transfer, lowering the pollution risks and saving working time;

•Avoiding cross contamination, sterilization by washing, and cleaning validation;

Powerful Platform Universality

•Suitable for culture of most suspension cells and flaky carrier;

•Compatible with culture modes including batch culture, batch feeding and perfusion;

Accurate Control System

•Accurate control of pendulum angle and rotation speed;

•Working with pH+DO sensors to offer the measurement of key process parameters;

•Omnidirectional monitoring of cell culture process;

Technical Parameters

| Item | 10L | 20L | 50L |
| --- | --- | --- | --- |
| Oscillator dimensions (mm) | 620*650*535 | 813*647*534 | 1075*806*627 |
| Oscillator weight (kg) | 60 | 65 | 100 |
| Control cabinet dimensions (mm) | 420*678*794 |  |  |
| Control cabinet weight(kg) | 50 |  |  |
| Power (KW) | 1.6 |  |  |
| Power requirement | AC 220V/50Hz |  |  |
| DO sensor | Range：0 - 250%，Accuracy： ±5% |  |  |
| pH sensor | Range：5.5 - 8.5，Accuracy：±0.1 |  |  |
| Temperature sensor | Range：-50-200℃, Accuracy：±(0.15+0.002 * [t]) |  |  |
| Load cell | Sensor Accuracy：C3 |  |  |
| Pump | 313D2*2,114DV*1 |  |  |
| Pressure sensor | Range：-10kpa-10kpa, Accuracy：±2%FS |  |  |
| Rocking | Rotation Speed Range：2 - 40rpm, Accuracy：±1rpm |  |  |
|  | Angle Range：2 - 12° , Accuracy：±0.5° |  |  |
| Temperature control | Control Range：RT- 42℃, Control Accuracy：±0.2℃ |  |  |
| pH control | Control Range：6.00-8.00，Control Accuracy：±0.05 |  |  |
| DO control | Control Range：0-100%，Control Accuracy：±3% |  |  |
| Exhaust gas heating control | Control Range：RT-70℃, Control Accuracy：±1℃ |  |  |

13

GER Wave-Mixed Bioreactor Bag

•Specially designed for common cell culture applications in biopharmaceuticals;

•Offering the appropriate environment for cell growth;

• Use the GER culture bag in conjunction with the GER reactor during the amplification process;

•The cultivation bag uses EB1596/9101 film material,

which is packaged and sterilized by gamma irradiation;

• Provide sheet-like carriers according to demand;

ISO9001:2015

Biological Safety

•The verified 9101 film and JYSS EB1596 film are suitable for cell culture；

• Bag integrity test, mitigating the use risks;

•Verified sterile gas filter ensures that the culture medium is not contaminated

Physical Properties

• USP<661.2>：Total organic carbon, absorbance pH value, appearance;

• USP<665>：Particulate matter;

• USP<778>：Extractables;

Physical Characteristics

•ASTM1003：Haze, clarity, light transmittance;

•ASTM882：Tensile strength, elongation at break, elasticity modulus;

•ASTM792：Density;

•ASTM3985：O2 permeability-<0.05cm³(m².d.bar);

•ASTM2576：CO2 permeability-<0.2cm³(m².d.bar);

•ASTM F1249：Vapor permeability-3.2g(m2.d);

Packaging Instructions

•TSE/BSE Statement;

• REACH Statement;

• Bisphenol A Free Statement;

•ASTM D4169:Package Transportation Verification;

Technical Parameters

A: Outlet 1/4" * 7/16" silicone tube (1000mm), straight joint, 1/4" * 7/16" TPE tube (500mm)

B,J: Needle free sampling valve

C: PH optical fiber electrode

D: DO optical fiber electrode

E: Inlet 1/4" * 7/16" silicone tube (1000mm), straight joint, 1/4" * 7/16" TPE tube (500mm)

F,G: Replenishing tube 1/8"*7/16" silicone tube (1000mm),straight joint, 1/8" *

1/4" welded tube (500mm)

Diagram of 50L GER wave-mixed bioreactor bag

H: Gas exhaust tube,1/4" * 7/16" TPE tube (65mm),Square filter,1/4"* 7/16" silicone tube (45mm),one-way valve

I: Air inlet tube 1/4" * 7/16" silicone tube (45mm),Circular filter

*Quality study and consumables’ quality standards comply with industrial standards and meet the requirements for drug declaration.

V8  2025.08